## Slide 1
Uncertainty Assessment for HATPRO Microwave Radiometer Measurements and Calibrations
Institute for Geophysics and Meteorology, University of Cologne | Tobias Böck, Bernhard Pospichal, Ulrich Löhnert
EMF Lecture Day, May 17th 2022
Session

## Slide 2: Motivation
The atmospheric boundary layer (ABL) is the most important under-sampled part of the atmosphere*
For NWP applications, the top-priority atmospheric variables like temperature (T) and humidity (H) profiles are not currently adequately measured**
T and H profiles can be obtained by ground-based microwave radiometer (MWR) observations
Yet MWR observations are not assimilated by any operational NWP system
*US National Research Council Reports **WMO guidance on observations for NWP: https://www.wmo.int/pages/prog/www/OSY/GOS-RRR.html
2

## Slide 3: Goals
Integration of ground-based thermodynamic profilers into the DWD forecasting system (TDYN-PRO)
EMF III: The potential of ground-based thermodynamic profilers for improving short-term weather forecasts
Q: Are the considered state-of-the-art ground-based profilers accurate and reliable enough for data assimilation at DWD?
Also important for E-PROFILE (EUMETNET)* and ACTRIS
*https://www.eumetnet.eu/activities/observations-programme/current-activities/e-profile/
3

## Slide 4: Instrument: HATPRO MWR
Humidity And Temperature PROfiler from Radiometer Physics GmbH
Measures thermal emissions from the atmos-phere in 14 different frequencies (H2O-band and O2-band)
Converts these radiances into Brightness Temperatures TBs [K] (through Planck’s law)
allows investigation of (via inversion):
    rudimentary humidity profiles
    temperature profiles
    integrated water vapor (IWV)
    liquid water path (LWP)
4
continuous data in all-sky conditions: resolution of seconds to minutes
commercially available for ~10 years, ready for network applications (in principle)

## Slide 5: HATPRO Calibration
For calibration, 2 blackbodies are necessary: Coldload and Hotload
Absolute calibration with liquid nitrogen (LN2), Coldload
  ~77.2 K boiling point
  should be done every 6 months
Relative calibration on Hotload every 5 min (automatic; target inside radiometer; a bit above ambient temp.)
5

## Slide 6: Accuracy of HATPROs (Instrument Errors)
Characterized
systematic errors:
  absolute calibration errors, biases (systematic differences between instruments
  drifts (instrument stability, TB differences before and after calibrations)
random errors:
  radiometric noise (e.g. via covariance matrices)
through coordinated experiments at JOYCE in Jülich and RAO in Lindenberg (FESSTVaL)
Assessed repeatability of calibrations
Objective: Define & apply procedure for MWR measurement uncertainty characterization for TBs (level1) per channel
 necessary for data assimilation applications
6

## Slide 7: Calibration Campaign at RAO in Lindenberg during FESSTVaL 2021
4 HATPROs on site (FOGHAT G5, DWDHAT G5, SUNHAT G2, HAMHAT G2)
Comparisons of zenith and blackbody measurements (to find out biases, drifts, noise levels, repeatability)
7
Thanks to Christine Knist for helping with the campaign and to the DWD Observatory in Lindenberg for providing the location and part of the data.

## Slide 8: Results Overview (instrument errors)
8

## Slide 9: Biases (Differences between newest model HATPROs) via Zenith Comparisons
9
Systematic differences
All errors are relative, there is no perfect absolute reference
ok
critical

## Slide 10: Repeatability: Differences at Zenith between Calibrations compared to a reference
10
Is influenced by the quality of the calibration
Can only be done this way when there are at least two MWRs on the same site simultaneously
Meaningful values can only be achieved with the same conditions for each calibration (same day, same weather, same target,…)
≤ 0.12 K
H2O-band
≤ 0.16 K
O2-band

## Slide 11: Drifts (via Coldload Differences between Calibrations)
11
Cannot be directly influenced by operator
Is also influenced by biases!
Need to wait a certain amount of time between calibrations to get meaningful data

## Slide 12: Short Reminder: Covariance Matrix
Correlated radiometric noise for all 14 channels (shows dependency of these channels)
The radiometric noise for a single channel can be determined by calculating the variance when looking on a stable blackbody target
Square root of diagonals is standard deviation
12

## Slide 13
13
Noise Levels for Coldload

## Slide 14
14
Noise Levels for Hotload
Coldload and Hotload Covariances similar for all Gen5 HATPROs, significantly higher values for older Gen2 instruments
Operator independent, instrument specific

## Slide 15
Results Overview (for Gen5 HATPROs):
15
Type of Error | Typical Error Values H2O-band | Typical Error Values O2-band | Determined via
Biases | usually ≤ 0.3 K (up to 0.48 K) | usually ≤ 0.5 K (up to 1.15 K) | Zenith measurement differences between two MWRs
Calibration Repeatability | ≤ 0.12 K | ≤ 0.16 K | Leaps to zenith reference measurements after two immediate consecutive calibrations
Drifts (over 6 months) | usually ≤ 0.3 K (up to 0.8 K) | usually ≤ 0.8 K (up to 1.3 K) | Differences at coldload after calibration
Noise Levels (coldload – hotload) (1s) | ≤ 0.11 K – 0.18 K | ≤ 0.27 K – 0.35 K | Standard deviation
Summary Instrument Errors
Comparison to RPG manual:
   absolute TB uncertainty 0.5 K,  noise 0.10 to 0.15 K
Suggestion for total maximum error per channel:
  Bias mean + Repeatability + Drift (for retrievals noise is also important)
For minimal error: only Repeatability

## Slide 16: Location dependent Errors: Radio Frequency Interference (RFI) and Obstacles
16

## Slide 17: MW-profiling: How does it work theoretically?
17
cloud contribution
Oxygen absorption  constant VMR, so only temperature dependent
O2-band channels are optically thick, especially ch 11-14:
MWR does not penetrate very far into the atmosphere
Measured TBs are Ts near the ground
pressure broadened WV line  total water vapor & weak profile information
WV channels are optically thin:
MWR can “see” far into the atmosphere
Is more susceptical to external errors like RFI and far away obstacles
K-Band
H2O vapor
(channel 1-7)
V-Band
O2
(channel 8-14)

## Slide 18: Obstacles and RFI at Jülich
18
30° Elevation Scan Interferences 14.04. – 18.05.2021
30° Elevation Scan Interferences 14.04. – 18.05.2020
Full 360° azimuth scans at 30° elevation for all 14 channels in clear-sky conditions
Probability of difference of: (azimuth step – azimuth minimum) < 1K
mast
RFI (communication links?)
Higher fluctuation of water vapor
courtesy of Johanna Drude

## Slide 19: Location of Research Center Jülich
19
WV and T fluctuations in different compass directions maybe due to this big open-cast mining pit near Jülich

## Slide 20: Impact of Obstacles on T-profiling
20
lower
higher
5°
30°
90°
optical thickness τ of atmosphere

## Slide 21: Sensitivity Study: Simulate Obstacles in Line of Sight
21
lower
higher
5°
30°
90°
optical thickness τ of atmosphere
distance x25°
distance x130°
distance x190°
distance x230°
distance x15°
Simulate obstacles by increasing optical thickness to very high values in a LBL-RT model

## Slide 22: Simulated Obstacles
22
For T-profiles, elevation scans are conducted in ch 11-14
Ch 11 should maybe not be used when an obstacle is nearby

## Slide 23: Outlook: Next Steps
Make retrievals (T-profiles) from TBs generated by the simulated obstacles and compare them with retrievals without obstacles
   paper?
Change lv1-files in incorporate biases, repeatability, drifts and noise levels and a total maximum and minimum error
Getting to know RTTOV-gb and use it with different elevation angles to create O-B statistics
Data Assimilation of MWR data into DWD forecast system
23

## Slide 24: Thank you for your attention!
24
Contact:
tobias.boeck@uni-koeln.de

## Slide 25: Bonus Slides
25

## Slide 26: FOGHAT at Köln: Azimuth Scans (Obstacles and external Interference)
26

## Slide 27: FOGHAT at Köln: Azimuth Scans (Obstacles and external Interference)
27
wall from building
between
NE and SE
EM interference?

## Slide 28: Biases/Offsets via Zenith Comparisons
28

## Slide 29: (Repeatability FOGHAT)
29